## Slide 1: Ultrasound Despeckling for Contrast Enhancement
Tay, P. C., Garson, C. D., Acton, S. T., & Hossack, J. A. (2010). Ultrasound despeckling for contrast enhancement. Image Processing, IEEE Transactions on, 19(7), 1847-1860.
Sonia H. Contreras Ortiz

## Slide 2: Introduction
Ultrasound is a widely used imaging modality in obstetrics and for the diagnosis and staging of a number of diseases.
Ultrasound is used as a reflection imaging modality that uses pulse waveforms with frequencies from 1 - 20MHz

## Slide 3: Introduction
Advantages:
  It is safe (does not use ionizing radiation).
  The transducer is small and easily manipulated.
  The image has enough resolution (0.2mm to 2mm) to display details of many structures of the body
  The imaging system is inexpensive, compact and mobile.
  Provides real-time images of blood velocity and flow

## Slide 4: Introduction
Limitations:
  Images are 2D, yet the anatomy is 3D, hence the diagnostician must integrate multiple images in his mind.
  The image represents a thin plane at some arbitrary angle in the body. It is difficult to localize the image plane and reproduce it at a later time for follow-up studies.
  Low image quality: speckle (“grainy” appearance), blurring, artifacts.
  Limited penetration, resolution is not isotropic.

## Slide 5: Introduction
The PSF represents the output of the ultrasound system to an ideal point target (impulse response)
PSF
0
1
2
x 10
-3
0
1
2
x 10
-3
PSF upscaled by 4
0
1
2
x 10
-3
0
1
2
x 10
-3
0
1
2
x 10
-3
0.2
0.4
0.6
0.8
1
Axial resolution
0
1
2
x 10
-3
0.2
0.4
0.6
0.8
1
Lateral resolution
0.2mm
0.64mm

## Slide 6: Introduction
Speckle results from the accumulation of random scatterings in the tissues.
Statistics of speckle vary depending on the number of scatterers per resolution cell.

## Slide 7: Introduction
Speckle models
  J(n,m): envelope detection amplitudes
  I(n,m): noise-free ideal image
  P(n,m): point spread function
  X(n,m): multiplicative speckle noise independent of I(n,m)
  +(n,m): additive speckle noise dependent of I(n,m)

## Slide 8: Background
Complex valued IQ data can be modeled as a sum of complex phasors for some positive integer K:
  Amplitudes in a constant reflectivity region are Rayleigh or Rician if K is large and and independent

## Slide 9: Background

## Slide 10: Background
Filtering techniques
  Adaptive filters
    Lee filter
    Frost filter
    Wiener filter …
  Anisotropic Diffusion
    Smoothes homogeneous regions while preserves edges.
Lee filter
Coefficient of variation

## Slide 11: Materials and methods
The proposed method consists of removing outliers aggressively (Adaptive parameter k is binary)
The variance in a homogeneous region is reduced and the mean value is maintained.
An outlier is defined as a local extremum.
The outlier is replaced by the local mean of the window (the outilier is not included).

## Slide 12: Materials and methods
Ideal
Noisy

## Slide 13: Results
Simulated images
  They used Field II (Ultrasound simulation based on Matlab)

## Slide 14: Results
a) Lee b) SRAD c) Wiener d) SBF (proposed)

## Slide 15: Results
In vivo mouse heart
Lee
SRAD

## Slide 16: Results
Wiener
SBF

## Slide 17: Results
SBF provided better segmentation in six out of eleven frames

## Slide 18: Conclusion
All tested despeckling algorithms provided robust segmentation of a mouse LV throughout the cardiac cycle.
However, SBF method consistently and more often provided contours that better resembled the manually defined contours.